OCEM E642 Power Supply – Complete Command Reference with Examples

Power Supply State & Control Commands

Command	Format	Description	Example

ON Regolatore	"ON<term>"	Turns PS ON (if in remote mode)	"ON<LF>"

OFF Regolatore	"OFF<term>"	Turns PS OFF	"OFF<LF>"

STANDBY Regolatore	"STB<term>"	Puts PS in STANDBY	"STB<LF>"

Polarità Positiva	"POS<term>"	Sets positive polarity	"POS<LF>"

Polarità Negativa	"NEG<term>"	Sets negative polarity	"NEG<LF>"

Circuito Aperto	"OPN<term>"	Opens current circuit	"OPN<LF>"

Reset Allarmi	"RES<term>"	Clears non-permanent alarms	"RES<LF>"

Start Rampa	"STR<term>"	Starts ramp execution	"STR<LF>"

Current & Ramp Control Commands

Command	Format	Description	Example

Set Point Corrente	"SP<space><value><term>"	Current setpoint (7 digits)	"SP 0500000<LF>"

Velocità Rampa Up	"VRUP<space><value><term>"	Upward ramp speed (7 digits)	"VRUP 0001000<LF>"

Velocità Rampa Down	"VRDW<space><value><term>"	Downward ramp speed (7 digits)	"VRDW 0001000<LF>"

Status Request Commands

Command	Format	Description	Response Example

Richiesta Stato Analogico	"SA<term>"	Current & voltage values	"COR 0500000<LF>TEN 0100000<LF>"

Richiesta Stato Logico	"SL<term>"	All status info	"SEL PRED<LF>STA ON<LF>POL POS<LF>ALL 01 03<LF>"

Richiesta Versione	"VER<term>"	Firmware version	"VER OCEM spa TeleoperatoreA950101PS_TYPE<LF>"

Configuration Commands

Programma Canale Analogico (PRG)

Format: "PRG<space><tipo><num><space><minval><space><maxval><space><um><term>"

Parameters:

<tipo>: I (input) or O (output)

<num>: Channel number (0-4)

<minval>: 7-digit minimum value

<maxval>: 7-digit maximum value

<um>: Unit of measure

Channel Mapping:

text

I0: Current value (16-bit digital)

I1: Voltage value (8-bit analog, 0-10V)

O0: Current setpoint (16-bit digital)

O1: Ramp up speed (12-bit analog, 0-10V)

O2: Ramp down speed (12-bit analog, 0-10V)

Examples:

ascii

"PRG I0 0000000 0500000 mA<LF>"    // Current: 0-500 mA

"PRG I1 0000000 0100000 mV<LF>"    // Voltage: 0-10 V

"PRG O1 0000000 0005000 AS<LF>"    // Ramp up: 0-5 A/s

"PRG O2 0000000 0003000 AS<LF>"    // Ramp down: 0-3 A/s

Soglia Analogica (TH)

Format: "TH<space>I<num><space><val><term>"

Sets threshold for spontaneous current/voltage messages.

Examples:

ascii

"TH I0 000500<LF>"    // Current threshold: 50 units

"TH I1 000100<LF>"    // Voltage threshold: 10 units

Channel Management Commands

Command	Format	Description	Example

Remote	"REM<term>"	Opens Primary Channel	"REM<LF>"

Local	"LOC<term>"	Opens Service Channel	"LOC<LF>"

Abilitazione	"ABL<term>"	Enables Service Channel	"ABL<LF>"

Disabilitazione	"DIL<term>"	Disables Service Channel	"DIL<LF>"

Spontaneous Messages (from PS)

Message	Format	Trigger	Example

Stato Regolatore	"STA<space><stato><term>"	State change	"STA ON<LF>"

Stato Selettore	"SEL<space><stato><term>"	Selector change	"SEL PRED<LF>"

Stato Polarità	"POL<space><stato><term>"	Polarity change	"POL POS<LF>"

Stato Allarmi	"ALL<list><term>"	Alarm change	"ALL 01 03 24<LF>"

Stato Valore Corrente	"COR<space><value><term>"	Current > threshold	"COR 0485000<LF>"

Stato Valore Tensione	"TEN<space><value><term>"	Voltage > threshold	"TEN 0098000<LF>"

Versione	"VER<space>OCEM spa Teleoperatore<vers><term>"	Power-on	"VER OCEM spa TeleoperatoreA950101PS_TYPE<LF>"

Cold Restart	"COLD RESTART<term>"	Lost configuration	"COLD RESTART<LF>"

Complete Operation Examples

1. Full PS Configuration Sequence

ascii

// 1. Enable remote control

"REM<LF>"

// 2. Configure analog channels

"PRG I0 0000000 0500000 mA<LF>"    // 0-500 mA current reading

"PRG O0 0000000 0500000 mA<LF>"    // 0-500 mA current setpoint

"PRG O1 0000000 0005000 AS<LF>"    // 0-5 A/s ramp up

"PRG O2 0000000 0003000 AS<LF>"    // 0-3 A/s ramp down

// 3. Set thresholds

"TH I0 000500<LF>"                  // 50 mA threshold

"TH I1 000100<LF>"                  // 1 V threshold

// 4. Turn ON and configure

"ON<LF>"

"POS<LF>"

"SP 0250000<LF>"                    // 250 mA setpoint

"VRUP 0002000<LF>"                  // 2 A/s ramp up

"VRDW 0001000<LF>"                  // 1 A/s ramp down

"STR<LF>"                           // Start ramp

2. Monitoring Sequence

ascii

// Request complete status

"SL<LF>"

// Expected response:

// SEL PRED<LF>      // Remote mode

// STA ON<LF>        // PS is ON

// POL POS<LF>       // Positive polarity

// ALL 01 03<LF>     // Alarms 01 and 03 active

// Request current values

"SA<LF>"

// Expected response:

// COR 0248000<LF>   // Current: 248 mA

// TEN 0095000<LF>   // Voltage: 9.5 V

3. Protocol-Specific Examples

Port A (RS485 - Polling-Selecting):

ascii

// Master polls Slave at address 1

Master: <ENQ>61

Slave:  <EOT>                         // No data

// Master sends command to Slave 1

Master: <ENQ>61

Slave:  <ACK>

Master: <STX>61SP 0250000<ETX><bcc>

Slave:  <ACK>

// Master reads response

Master: <ENQ>61

Slave:  <STX>61STA ON<ETX><bcc>

Master: <ACK>

Slave:  <EOT>

Port B (RS232 - Simple ASCII):

ascii

// Direct command and response

Host: "SP 0250000<LF>"

PS:   "COR 0248000<LF>"               // Response after execution

Important Protocol Notes

Addressing (Port A only):

Slave addresses: 0-31

Polling address = 0x60 + slave address

Selecting address = 0x60 + slave address

Numeric Format:

All values are 7-digit integers with leading zeros

Example: 500 mA = "0500000"

Checksum Calculation (Port A):

python

# Example for "STA ON" to slave address 1

message = b"\x02\x61STA ON\x03"  # <STX><addr>STA ON<ETX>

bcc = 0

for byte in message[1:]:  # Skip <STX>

bcc ^= byte

bcc |= 0x80  # Set MSB to 1

Execution Conditions:

State commands require selector in "PRED" (remote) mode

Commands accepted in any state but executed only if compatible

Default scaling applied if PRG not configured

This reference includes all PS operations and protocol commands with practical examples. Use Port A for reliable multi-device control and Port B for local diagnostics only.